LA-ICPMS methods

Zircon grains were separated from rocks using standard techniques, annealed at 900oC for 60 hours in a muffle furnace, and randomly selected grains were mounted in epoxy and polished until their centers were exposed. Cathodoluminescence (CL) images were obtained with a JEOL JSM-300 scanning electron microscope and Gatan MiniCL. Zircon was analyzed by laser ablation inductively coupled plasma mass spectrometry (LA-ICPMS) using a ThermoElectron X-Series II quadrupole ICPMS and New Wave Research UP-213 Nd:YAG UV (213 nm) laser ablation system. In-house analytical protocols, standard materials, and data reduction software were used for acquisition and calibration of U-Pb dates and a suite of high field strength elements (HFSE) and rare earth elements (REE). Zircon was ablated with a laser spot of 25 µm wide using fluence and pulse rates of 5 J/cm2 and 5 Hz, respectively, during a 45 second analysis (15 sec gas blank, 30 sec ablation) that excavated a pit ~25 µm deep. Ablated material was carried by a 1.2 L/min He gas stream to the nebulizer flow of the plasma. Quadrupole dwell times were 5 ms for Si and Zr, 200 ms for 49Ti and 207Pb, 80 ms for 206Pb, 40 ms for 202Hg, 204Pb, 208Pb, 232Th, and 238U and 10 ms for all other HFSE and REE; total sweep duration is 950 ms. Background count rates for each analyte were obtained prior to each spot analysis and subtracted from the raw count rate for each analyte. For concentration calculations, background-subtracted count rates for each analyte were internally normalized to 29Si and calibrated with respect to NIST SRM-610 and -612 glasses as the primary standards. Ablations pits that appear to have intersected glass or mineral inclusions were identified based on Ti and P signal excursions, and associated sweeps were discarded. U-Pb dates from these analyses are considered valid if the U-Pb ratios appear to have been unaffected by the inclusions. Signals at mass 204 were normally indistinguishable from zero following subtraction of mercury backgrounds measured during the gas blank (<1000 cps 202Hg), and thus dates are reported without common Pb correction. Rare analyses that appear contaminated by common Pb were rejected based on mass 204 greater than baseline. Temperature was calculated from the Ti-in-zircon thermometer (Watson et al., 2006). Because there are no constraints on the activity of TiO2, an average value in crustal rocks of 0.8 was used.

Data were collected in one experiment in May 2017. For U-Pb and 207Pb/206Pb dates, instrumental fractionation of the background-subtracted ratios was corrected and dates were calibrated with respect to interspersed measurements of zircon standards and reference materials. The primary standard Plešovice zircon (Sláma et al., 2008) was used to monitor time-dependent instrumental fractionation based on two analyses for every 10 analyses of unknown zircon. A secondary correction to the 206Pb/238U dates was made based on results from the zircon standards Seiland (531 Ma, unpublished data, Boise State University) and Zirconia (327 Ma, unpublished data, Boise State University), which were treated as unknowns and measured once for every 10 analyses of unknown zircon. These results showed a linear age bias of several percent that is related to the 206Pb count rate. The secondary correction is thought to mitigate matrix-dependent variations due to contrasting compositions and ablation characteristics between the Plešovice zircon and other standards (and unknowns).

Radiogenic isotope ratio and age error propagation for all analyses includes uncertainty contributions from counting statistics and background subtraction. The standard calibration uncertainty for U/Pb is the local standard deviation of the polynomial fit to the fractionation factor of Plešovice versus time and for 207Pb/206Pb is the standard error of the mean of the fractionation factor of Plešovice. These uncertainties are 1.2% (2) for 206Pb/238U and 1.1% (2) for 207Pb/206Pb. Age interpretations are based on 206Pb/238U dates. Errors are at 2.

## CA-TIMS U-Pb Geochronology Methods

Zircon was removed from the epoxy mounts for dating based on CL images and LA-ICPMS data. Grains were put into 3 ml Teflon PFA beakers and loaded into 300 l Teflon PFA microcapsules. Fifteen microcapsules were placed in a large-capacity Parr vessel and the zircon partially dissolved in 120 l of 29 M HF for 12 hours at 190°C. Zircon was returned to 3 ml Teflon PFA beakers, HF was removed, and zircon was immersed in 3.5 M HNO3, ultrasonically cleaned for an hour, and fluxed on a hotplate at 80°C for an hour. The HNO3 was removed and zircon was rinsed twice in ultrapure H2O before being reloaded into the 300 l Teflon PFA microcapsules (rinsed and fluxed in 6 M HCl during sonication and washing of the zircon) and spiked with the Boise State University mixed 233U-235U-205Pb tracer solution (BSU-1B). Zircon was dissolved in Parr vessels in 120 l of 29 M HF with a trace of 3.5 M HNO3 at 220°C for 48 hours, dried to fluorides, and re-dissolved in 6 M HCl at 180°C overnight. U and Pb were separated from the zircon matrix using an HCl-based anion-exchange chromatographic procedure (Krogh, 1973), eluted together and dried with 2 µl of 0.05 N H3PO4.

Pb and U were loaded on a single outgassed Re filament in 5 µl of a silica-gel/phosphoric acid mixture (Gerstenberger and Haase, 1997), and U and Pb isotopic measurements made on a GV Isoprobe-T multicollector thermal ionization mass spectrometer equipped with an ion-counting Daly detector. Pb isotopes were measured by peak-jumping all isotopes on the Daly detector for 160 cycles, and corrected for 0.16 ± 0.03%/a.m.u. (1) mass fractionation. Transitory isobaric interferences due to high-molecular weight organics, particularly on 204Pb and 207Pb, disappeared within approximately 60 cycles, while ionization efficiency averaged 104 cps/pg of each Pb isotope. Linearity (to ≥1.4 x 106 cps) and the associated deadtime correction of the Daly detector were determined by analysis of NBS982. Uranium was analyzed as UO2+ ions in static Faraday mode on 1012 ohm resistors for 300 cycles, and corrected for isobaric interference of 233U18O16O on 235U16O16O with an 18O/16O of 0.00206. Ionization efficiency averaged 20 mV/ng of each U isotope. U mass fractionation was corrected using the known 233U/235U ratio of the Boise State University tracer solution.

U-Pb dates and uncertainties were calculated using the algorithms of Schmitz and Schoene (2007), calibration of BSU-1B tracer solution of 235U/205Pb of 77.93 and 233U/235U of 1.007066 for, U decay constants recommended by Jaffey et al. (1971), and 238U/235U of 137.818 (Hiess et al., 2012). 206Pb/238U ratios and dates were corrected for initial 230Th disequilibrium using DTh/U = 0.20 ± 0.05 (1) and the algorithms of Crowley et al. (2007), resulting in an increase in the 206Pb/238U dates of ~0.09 Ma. All common Pb in analyses was attributed to laboratory blank and subtracted based on the measured laboratory Pb isotopic composition and associated uncertainty. U blanks are estimated at 0.013 pg.

A weighted mean 206Pb/238U date are calculated from equivalent dates (probability of fit >0.05) using Isoplot 3.0 (Ludwig, 2003). Errors on weighted mean date are given as ± x / y / z, where x is the internal error based on analytical uncertainties only, including counting statistics, subtraction of tracer solution, and blank and initial common Pb subtraction, y includes the tracer calibration uncertainty propagated in quadrature, and z includes the 238U decay constant uncertainty propagated in quadrature. Internal error should be considered when comparing our dates with 206Pb/238U dates from other laboratories that used the same tracer solution or a tracer solution that was cross-calibrated using EARTHTIME gravimetric standards. Error including the uncertainty in the tracer calibration should be considered when comparing our date with those derived from other geochronological methods using the U-Pb decay scheme (e.g., laser ablation ICPMS). Error including uncertainties in the tracer calibration and 238U decay constant (Jaffey et al., 1971) should be considered when comparing our date with those derived from other decay schemes (e.g., 40Ar/39Ar, 187Re-187Os). Errors are at 2.

## U-Pb Geochronology Results

Thirty-one zircon grains from 16SI166 analyzed by LA-ICPMS yield dates from 56 ± 3 to 47 ± 3 Ma. Eight grains were analyzed by CA-TIMS. The four youngest dates yield a weighted mean of 50.906 ± 0.019 (MSWD = 2.2, probability of fit = 0.09). This is the interpreted igneous crystallization age. Four other grains that yield dates of 51.043 ± 0.041 to 50.954 ± 0.040 Ma are interpreted as containing inherited components.

## References

Crowley, J.L., Schoene, B., Bowring, S.A., 2007, U-Pb dating of zircon in the Bishop Tuff at the millennial scale: Geology 35: 1123-1126.

Gerstenberger, H., Haase, G., 1997, A highly effective emitter substance for mass spectrometric Pb isotope ratio determinations: Chemical Geology 136: 309-312.

Hiess, J., Condon, D.J., McLean, N., Noble, S. R., 2012, 238U/235U systematics in terrestrial uranium-bearing minerals: Science 335: 1610-1614.

Jaffey, A.H., Flynn, K.F., Glendenin, L.E., Bentley, W.C., Essling, A.M., 1971, Precision measurements of half-lives and specific activities of 235U and 238U: Physical Review C, 4: 1889-1906.

Krogh, T.E., 1973, A low contamination method for hydrothermal decomposition of zircon and extraction of U and Pb for isotopic age determination: Geochimica et Cosmochimica Acta 37: 485-494.

Ludwig, K.R., 2003, User’s Manual for Isoplot 3.00. Berkeley Geochronology Center: Berkeley, CA, 70 p.

Mattinson, J.M., 2005, Zircon U-Pb chemical abrasion ("CA-TIMS") method: combined annealing and multi-step partial dissolution analysis for improved precision and accuracy of zircon ages: Chemical Geology 220:47-66.

Schmitz, M.D., Schoene, B., 2007, Derivation of isotope ratios, errors and error correlations for U-Pb geochronology using 205Pb-235U-(233U)-spiked isotope dilution thermal ionization mass spectrometric data: Geochemistry, Geophysics, Geosystems (G3) 8, Q08006, doi:10.1029/2006GC001492.

Sláma, J., Košler, J, Condon, D.J., Crowley, J.L., Gerdes, A., Hanchar, J.M., Horstwood, M.S.A., Morris, G.A., Nasdala, L., Norberg, N., Schaltegger, U., Schoene, B. Tubrett, M.N, Whitehouse, M.J., 2008, Plešovice zircon — A new natural reference material for U–Pb and Hf isotopic microanalysis. Chemical Geology, 249: 1-35.

Watson, E.B., Wark, D.A., Thomas, J.B., 2006, Crystallization themometers for zircon and rutile. Contributions to Mineralogy and Petrology, 151: 413-433.

Figure captions

Plot of CA-TIMS 206Pb/238U dates from zircon. Plotted with Isoplot 3.0 (Ludwig, 2003). Error bars are at 2 sigma. Weighted mean date is shown and represented by gray box behind the error bars. MSWD = Mean Square of Weighted Deviation. pof = probability of fit.

CL images of zircon. LA-ICPMS spot locations and labels are shown in red. Labels for grains dated by CA-TIMS are shown in yellow.